## Slide 1: DNN-Assisted CLL MRD Detection by Flow Cytometry Clinical validation and post-implementation monitoring
Published : 2025
Location : Mayo Clinic, USA;

## Slide 2: Journal Metrics
Journal Name | Cancers
Publisher | MDPI
Impact Factor | 4.4
Cite Score | 8.8
Journal Rank | Q1

## Slide 3: Agenda
Introduction
Methods
Results
Discussion

## Slide 4: Introduction

## Slide 5: Overview

## Slide 6: Overview

## Slide 7: Why MRD in CLL matters
MRD status is a powerful predictor of outcomes in CLL and increasingly guides therapy decisions.
Flow cytometric MRD detection is standard but is complex, time‑consuming, and operator‑dependent.

## Slide 8: Study Aim
Evaluate a deep neural network (DNN)–assisted, human‑in‑the‑loop workflow for CLL MRD detection in a clinical lab.
Perform rigorous clinical validation and implement continuous post‑deployment performance monitoring.

## Slide 9: Figure 1: Comparison of manual and ai-assisted workflows for flowcytometry analysis in cll mrd.

## Slide 10: Methods

## Slide 11: Methods — Method comparison
Design: DNN + 2nd review versus manual reference method on 240 consecutive clinical samples.
Sample composition: ≥30% MRD‑positive and ≥30% MRD‑negative; MRD‑positive span full range.
Quantitative comparison via Deming regression; qualitative metrics included sensitivity/specificity.

## Slide 12: Methods — Precision verification
Quantitative: Six MRD‑positive samples (<1%) re‑analyzed in triplicate across two instruments and two operators (12 replicates/sample).
Qualitative: 14 historical cases (10 positive, 4 negative) × 4 replicates each, analyzed by 4 technologists in a blinded fashion.

## Slide 13: Methods — Analytical sensitivity
Assay LLOQ: 0.002% (20 clonal events per 1,000,000 WBCs); LOD: 0.001%.
Verification used reanalysis of prior dilution series at 0.02%, 0.002% (triplicates), and 0.001% (triplicates).

## Slide 14: Post‑implementation monitoring
Daily electronic QC: hash‑based verification of full pipeline on a control file for each workstation/user.
Input data drift detection: PCA reconstruction error thresholds; alerts logged and visualized.
Error analysis monitoring: event‑level TP/TN/FP/FN versus human‑corrected labels, tracked over time.
Attribute acceptance sampling: periodic expert re‑review of random 'negative' reports to detect false negatives.

## Slide 15: Results

## Slide 16: Results — Binary classification accuracy
Overall accuracy: 97.5% (234/240) versus manual reference.
Sensitivity: 100% (129/129);
Specificity: 94.6% (105/111).
Performance consistent across bone marrow (96.1%) and peripheral blood (97.9%).

## Slide 17: Results — Quantitative agreement
Correlation r = 0.99 and slope = 0.99 between DNN + 2nd review and manual reference.
No significant difference between bone marrow (51) and peripheral blood (189) slopes/intercepts.

## Slide 18: DNN‑only vs human‑in‑the‑loop
DNN‑only correlation r ≈ 0.97; slope ≈ 0.95;
overall accuracy ≈ 90.4% with NPV ≈ 94%.
Missed 13/874 positives (1.5%) later recovered by technologist review; many had atypical or non‑CLL phenotypes.

## Slide 19: Figure 2
Notes: Figure 2. Correlation between manual analysis (ground truth) and DNN-based approaches for quantitative MRD assessment. (A) Scatter plot showing the correlation between the manual reference method and DNN + 2nd review method for the percentage of clonal events per analyzed white blood cells (WBCs), with a correlation coefficient of 0.99 and slope of 0.99. (B) Scatter plot showing the correlation between ground truth manual analysis and DNN-only analysis. The Deming regression shows a correlation coefficient of 0.970 and slope of 0.95. The red dot highlights an outlier case with mantle cell lymphoma that was misclassified by the DNN (B), and was subsequently corrected by the expert upon 2nd review (A).

## Slide 20: Results — Precision
Quantitative precision met verification limits for 5/6 samples; one (BM1) slightly above repeatability SD limit but within within‑lab SD.
Qualitative precision: no significant differences across four technologists; near‑perfect accordance/concordance.

## Slide 21: Table 1
Sample | Mean (MRD %) | Repeatability SD | Repeatability CV | Within Laboratory SD | Within Laboratory CV | Upper Verification Limit—Repeatability SD | Upper Verification Limit—Within Laboratory SD
BM1 | 0.75262 | 0.06798 | 9.00% | 0.07097 | 9.40% | 0.06636 | 0.16432
BM2 | 0.37903 | 0.01444 | 3.80% | 0.02269 | 6.00% | 0.01919 | 0.03510
BM3 | 0.36078 | 0.01682 | 4.70% | 0.01906 | 5.30% | 0.02735 | 0.02528
BM4 | 0.55410 | 0.03104 | 5.60% | 0.03104 | 5.60% | 0.04622 | 0.06913
BM5 | 0.67446 | 0.05677 | 8.40% | 0.05926 | 8.80% | 0.06911 | 0.11944
PB1 | 0.02114 | 0.00337 | 16.00% | 0.00399 | 18.90% | 0.00667 | 0.00616

## Slide 22: Results — Analytical sensitivity
All three samples detected at 0.002% with intra‑assay CVs 6.6%–23.5% (all <30%).
Assay sensitivity meets/exceeds 10⁻4–10⁻5 levels recommended for CLL MRD.

## Slide 23: Results — Drift detection in practice
Over 6 months: 874 samples analyzed; 26 (2.97%) triggered drift alerts; no day met formal drift event criteria.
Most alerts were high‑burden CLL (>40% abnormal cells); others included MCL, CD5‑negative disorders, and atypical CLL.

## Slide 24: Table 2
Alert Category | Count | Percentage | Drift Value Range
High-burden CLL (>40%) | 17 | 65.4% | 1.85–3.99
Mantle cell lymphoma | 3 | 11.5% | 1.78–2.51
CD5-negative B-cell disorders | 4 | 15.4% | 1.92–2.95
Atypical CLL immunophenotype | 2 | 7.7% | 1.79–1.84

## Slide 25: Figure 3

## Slide 26: Results — Error analysis monitoring
DNN‑only vs corrected labels:
sensitivity ≈97%,
specificity ≈59%,
PPV ≈73%,
NPV ≈95%.
Pattern of misses informs training gaps and triage rules for human review.

## Slide 27: Table 3
 | Number (% of 874 Cases)
Missed cases by DNN-only | 13 (1.5%)
CD5-negative | 5 (0.6%)
CD5-positive | 8 (0.9%)
Non-CLL | 7
CD20 moderate | 7
CD22 moderate | 7
CD43-negative or partial | 7
CD45 not dim | 7
Light chain moderate | 7
CD200-negative | 4
CLL | 1 *

## Slide 28: Results — Attribute acceptance sampling
Semiannual re‑review of 84 negative cases found 98.8% were true negatives (1/84 false negative).
Meets acceptance criterion (≤1 failure), supporting reliability of negative results in routine use.

## Slide 29: Results — Time savings and efficiency
Average analysis time reduced from 10.5 ± 5.8 to 4.2 ± 2.3 minutes (≈60.3% reduction).
Estimated ~158 technologist hours saved per year for ~1500 cases, enabling capacity and turnaround gains.

## Slide 30: Table 4
Analysis Method | Mean Time (min) | Standard Deviation (min) | Range (min) | Time Reduction
Pre-implementation manual analysis | 10.5 | 5.8 | 3.5–27.5 | —
Initial DNN-assisted workflow (go-live) | 3.9 | 2.1 | 1.0–9.0 | 62.9%
Current DNN-assisted workflow (1-year follow-up) | 4.2 | 2.3 | 1.0–10.5 | 60.3%

## Slide 31: Discussion

## Slide 32: Clinical implications
Human‑in‑the‑loop AI achieves expert‑level accuracy while accelerating throughput.
Monitoring framework provides safety net for drift and atypical phenotypes, maintaining clinical reliability.

## Slide 33: limitation
Single institution and specific 10‑color panel; generalizability requires local training/validation.
Model focused on CLL; atypical/non‑CLL immunophenotypes remain challenging without human review.

## Slide 34: Future direction
Improve specificity without sacrificing sensitivity (e.g., class weighting, focal loss, additional markers).
Advance monitoring with finer drift attribution and add uncertainty quantification to auto‑flag edge cases.

## Slide 35: Messages
DNN‑assisted, human‑vetted CLL MRD analysis matches expert performance and cuts analysis time by ~60%.
Continuous monitoring (QC, drift, error analysis, and acceptance sampling) keeps performance safe over time.

## Slide 36: Thanks for your attention
Milad Rahimi
rahimimiladofficial@gmail.com